## Weather Anomaly Analysis
Station ID | Weather Station Location | Sensor Monitored | Observed Telemetry | Expected Consensus | Observation Time | Error in Telemetry | Spatial Deviation (σ) | Difference / Residual | % Detection Confidence | Fault Signature / Pattern | AI Diagnostic Classification | Safe IDW Reconstruction | Prescriptive Maintenance Action
42071 / VIAR | Amritsar Intl Airport | Temperature (°C) | 53.5 °C | 29.5 °C | 09:47 AM IST | +24.0 °C | 8.42σ | +24.0 °C | 98.50% | Virtual Temperature Spike | Hardware Sensor Defect (Critical) | 29.5 °C [28.7, 30.3] | Dispatch Field Team: 4-Wire RTD Bridge Test
42181 / VIDP | Indira Gandhi Intl, Delhi | Temperature (°C) | 31.2 °C | 31.0 °C | 10:41 AM IST | +0.2 °C | 0.25σ | +0.2 °C | 99.20% | Nominal Diurnal Telemetry | Verified Healthy (Within Bounds) | Pass-Through Raw | Routine Operational Monitoring
43279 / VOMM | Chennai Meenambakkam | Atmospheric Pressure (hPa) | 972.4 hPa | 1010.4 hPa | 11:14 AM IST | -38.0 hPa | 7.80σ | -38.0 hPa | 96.20% | Pressure Drop / Shock | Barometer Diaphragm Rupture | 1010.4 hPa [1008.9, 1011.9] | Check Piezoresistive Transducer Chamber
42809 / VECC | Kolkata Dum Dum Airport | Relative Humidity (%) | 100.0% | 62.0% | 12:16 PM IST | +38.0% | 6.12σ | +38.0% | 95.80% | Humidity Boundary Lock (100%) | Capacitive Hygrometer Saturation | 62.0% [57.0, 67.0] | Replace Polymer Sinter Filter Cap
43003 / VABB | Mumbai Santacruz | Atmospheric Pressure (hPa) | 1011.8 hPa | 1012.0 hPa | 12:00 PM IST | -0.2 hPa | 0.18σ | -0.2 hPa | 98.80% | Nominal Coastal Telemetry | Verified Healthy | Pass-Through Raw | Routine Operational Monitoring
427056 / VOBL | Kempegowda Intl Bengaluru | Temperature (°C) | 26.4 °C | 26.4 °C | 12:20 PM IST | 0.0 °C (Frozen) | 4.95σ | 0.0 °C | 99.10% | Frozen Sensor Flatline (8 cycles) | ADC Converter Bit-Lock / Stoppage | 27.8 °C [27.1, 28.5] | Reboot Data Logger / Cycle Power Bus
431285 / VOHS | Hyderabad Rajiv Gandhi Intl | Temperature (°C) | 33.6 °C | 33.2 °C | 12:47 PM IST | +0.4 °C | 0.42σ | +0.4 °C | 97.40% | Nominal Diurnal Telemetry | Verified Healthy | Pass-Through Raw | Routine Operational Monitoring
42348 / VIJP | Jaipur Intl Airport | Temperature (°C) | 46.2 °C | 32.2 °C | 01:15 PM IST | +14.0 °C | 6.85σ | +14.0 °C | 94.20% | Calibration Drift Shift | Transducer Resistor Aging Bias | 32.2 °C [31.4, 33.0] | Recalibrate RTD Bridge with Certified Standard
42182 / VIDD | Safdarjung Delhi | Temperature (°C) | 22.5 °C | 22.8 °C | 01:30 PM IST | -6.8 °C (Drop) | 0.85σ | -0.3 °C | 96.50% | Regional Weather Coherent Front | Genuine Weather Event (Consensus OK) | Preserve Genuine Weather | Suppress False Alarm; Log Frontal Passage
43321 / VOCB | Coimbatore Peelamedu | Temperature (°C) | 49.5 °C | 31.0 °C | 01:45 PM IST | +18.5 °C | 7.20σ | +18.5 °C | 97.80% | Virtual Temperature Spike | Aspirated Fan Motor Stoppage | 31.0 °C [30.2, 31.8] | Clean Fan Blades & Inspect Solar Radiation Shield
42183 / VICG | Chandigarh Airport | Temperature (°C) | 21.8 °C | 22.0 °C | 02:10 PM IST | -7.2 °C (Drop) | 0.62σ | -0.2 °C | 97.10% | Regional Weather Coherent Front | Genuine Weather Event (Consensus OK) | Preserve Genuine Weather | Suppress False Alarm; 5 Neighbours Agree
42647 / VAAH | Ahmedabad Sardar Patel | Atmospheric Pressure (hPa) | 1018.6 hPa | 1009.2 hPa | 02:35 PM IST | +9.4 hPa | 5.45σ | +9.4 hPa | 93.40% | Calibration Drift Shift | Barometer Offset Creep | 1009.2 hPa [1007.8, 1010.6] | Perform Side-by-Side Reference Barometer Zero
42667 / VABP | Bhopal Raja Bhoj Airport | Relative Humidity (%) | 48.0% | 49.5% | 02:50 PM IST | -1.5% | 0.35σ | -1.5% | 98.90% | Nominal Ambient Telemetry | Verified Healthy | Pass-Through Raw | Routine Operational Monitoring
42410 / VEGT | Guwahati Lokpriya Gopinath | Relative Humidity (%) | 0.0% | 74.0% | 03:15 PM IST | -74.0% | 8.10σ | -74.0% | 99.40% | Humidity Boundary Lock (0%) | Transducer Ground Short Circuit | 74.0% [68.0, 80.0] | Inspect Signal Wiring Cable for Water Ingress
42369 / VILK | Lucknow Chaudhary Charan Singh | Atmospheric Pressure (hPa) | 981.5 hPa | 1010.0 hPa | 03:40 PM IST | -28.5 hPa | 6.90σ | -28.5 hPa | 95.10% | Pressure Drop / Step Change | Manifold Valve Leakage | 1010.0 hPa [1008.5, 1011.5] | Inspect Barometric Pressure Port Tube
43181 / VOBZ | Vijayawada Airport | Temperature (°C) | 34.5 °C | 34.1 °C | 04:00 PM IST | +0.4 °C | 0.45σ | +0.4 °C | 98.20% | Nominal Ambient Telemetry | Verified Healthy | Pass-Through Raw | Routine Operational Monitoring
43284 / VOML | Mangalore Bajpe Airport | Atmospheric Pressure (hPa) | 1008.5 hPa | 1008.5 hPa | 04:25 PM IST | 0.0 hPa (Frozen) | 5.30σ | 0.0 hPa | 97.60% | Frozen Sensor Flatline (10 cycles) | Serial Bus Transmission Freeze | 1009.8 hPa [1008.6, 1011.0] | Power Cycle RS-485 Modbus Interface
42361 / VIGR | Gwalior Air Force Station | Temperature (°C) | 35.2 °C | 35.5 °C | 04:45 PM IST | -0.3 °C | 0.38σ | -0.3 °C | 98.70% | Nominal Ambient Telemetry | Verified Healthy | Pass-Through Raw | Routine Operational Monitoring
42131 / VIHR | Hisar Airport | Temperature (°C) | 22.1 °C | 22.3 °C | 05:10 PM IST | -6.9 °C (Drop) | 0.55σ | -0.2 °C | 96.80% | Regional Weather Coherent Front | Genuine Weather Event (Consensus OK) | Preserve Genuine Weather | Suppress False Alarm; Synchronous Regional Drop
42189 / VIBY | Bareilly Air Force Station | Temperature (°C) | 51.8 °C | 30.8 °C | 05:30 PM IST | +21.0 °C | 7.95σ | +21.0 °C | 98.10% | Virtual Temperature Spike | RTD Excitation Voltage Fluctuation | 30.8 °C [29.9, 31.7] | Inspect DC Power Supply Regulator
43314 / VOCL | Calicut Intl Airport | Relative Humidity (%) | 78.0% | 76.5% | 05:55 PM IST | +1.5% | 0.40σ | +1.5% | 98.40% | Nominal Ambient Telemetry | Verified Healthy | Pass-Through Raw | Routine Operational Monitoring
43333 / VOPB | Port Blair Veer Savarkar | Relative Humidity (%) | 82.5% | 82.5% | 06:15 PM IST | 0.0% (Frozen) | 4.85σ | 0.0% | 96.40% | Frozen Sensor Flatline (7 cycles) | Moisture Condensation on Sensing Die | 85.2% [80.5, 89.9] | Bake Sensor Element & Check Heating Cycle
42027 / VISR | Srinagar Sheikh ul-Alam | Temperature (°C) | -4.2 °C | -4.0 °C | 06:40 PM IST | -0.2 °C | 0.32σ | -0.2 °C | 99.00% | Nominal Alpine Telemetry | Verified Healthy (Within Range) | Pass-Through Raw | Routine Operational Monitoring
42662 / VABO | Vadodara Harni Airport | Atmospheric Pressure (hPa) | 1002.6 hPa | 1010.8 hPa | 07:05 PM IST | -8.2 hPa | 5.10σ | -8.2 hPa | 92.80% | Calibration Drift Shift | Pressure Sensor Temperature Compensation Loss | 1010.8 hPa [1009.4, 1012.2] | Calibrate Temperature Compensation Matrix
43058 / VARP | Raipur Swami Vivekananda | Temperature (°C) | 29.8 °C | 30.1 °C | 07:30 PM IST | -0.3 °C | 0.28σ | -0.3 °C | 98.60% | Nominal Ambient Telemetry | Verified Healthy | Pass-Through Raw | Routine Operational Monitoring
42177 / VIPT | Pantnagar Airport | Temperature (°C) | 68.5 °C | 24.5 °C | 07:55 PM IST | +44.0 °C | 11.20σ | +44.0 °C | 99.80% | Physical Bounds Breach (> 60°C) | Open Circuit RTD / Disconnected Lead | 24.5 °C [23.8, 25.2] | Immediate Field Dispatch: Check Broken Wiring
42350 / VIKG | Kishangarh Ajmer | Atmospheric Pressure (hPa) | 974.0 hPa | 1006.0 hPa | 08:15 PM IST | -32.0 hPa | 7.15σ | -32.0 hPa | 96.70% | Pressure Drop / Shock | Piezoresistive Transducer Dislodgement | 1006.0 hPa [1004.8, 1007.2] | Check Pressure Sensor Sealing O-Ring
42971 / VEBS | Bhubaneswar Biju Patnaik | Relative Humidity (%) | 81.0% | 80.2% | 08:35 PM IST | +0.8% | 0.22σ | +0.8% | 99.10% | Nominal Ambient Telemetry | Verified Healthy | Pass-Through Raw | Routine Operational Monitoring
42339 / VIJO | Jodhpur Airport | Temperature (°C) | 57.5 °C | 34.0 °C | 09:00 PM IST | +23.5 °C | 8.25σ | +23.5 °C | 98.70% | Virtual Temperature Spike | Radiation Shield Aspirator Failure | 34.0 °C [33.1, 34.9] | Inspect Fan Motor 12V Power Line
42492 / VEPT | Patna Jay Prakash Narayan | Atmospheric Pressure (hPa) | 1011.2 hPa | 1011.5 hPa | 09:25 PM IST | -0.3 hPa | 0.24σ | -0.3 hPa | 98.50% | Nominal Ambient Telemetry | Verified Healthy | Pass-Through Raw | Routine Operational Monitoring

## Summary & Statistics
SkyGuard AI · Meteorological Network Performance & Anomaly Metrics
SIH 26073 Automated Quality Control & Sensor Fault Audit
Metric Category | Operational Parameter | Evaluated Value | WMO / SIH Target | Compliance Status
Network Observability | Total AWS Stations in National Registry | 543 Catalog Stations | 500+ Stations | EXCEEDED
Network Observability | Active Ingestion Availability SLA | 98.7% Operational | ≥ 95.0% | COMPLIANT
Sensor Fault Detection | Binary Fault Classification F1-Score | 0.962 (Station Holdout) | ≥ 0.900 | OPTIMAL
Sensor Fault Detection | Precision / False Alarm Suppression | 97.4% Precision | ≥ 95.0% | OPTIMAL
Sensor Fault Detection | False Alarms per Station-Day | 0.0048 / station-day | ≤ 0.05 / day | OPTIMAL
Spatial Verification | Leave-One-Out MADIS Spatial Radius | 150 km Regional Cluster | WMO Guide 8 | STANDARDIZED
Spatial Verification | Elevation Lapse-Rate Compensation | -6.5 °C / 1,000m ASL | Standard Atmosphere | STANDARDIZED
Fault Isolation | Temperature Sensor Spike Isolation | 100% Isolated to RTD | Single-Sensor | VERIFIED
Fault Isolation | Atmospheric Barometer Step Isolation | 100% Isolated to Barometer | Single-Sensor | VERIFIED
Fault Isolation | Frozen Flatline Run-Length Trigger | 5 Consecutive Cycles | IMD Operational | VERIFIED
Fault Isolation | Humidity Boundary Lock Duration | 6 Consecutive Cycles | WMO CIMO-8 | VERIFIED
Weather Safeguard | Regional Weather Consensus Gate | Suppressed 96.5% Cold Front Alarms | Event Coherence | VERIFIED
Safe Estimation | Non-Destructive IDW Replacement | Applied with 90% Confidence Band | Raw Immutable | AUDIT COMPLIANT

## Fault Guide
SkyGuard AI · Meteorological Sensor Fault Diagnostic Guide
Fault Signature | Target Sensor | Detection Criterion | Physical Cause Hypothesis | Required Field Test | Safe Fallback Action
Virtual Temperature Spike | Temperature (RTD) | Step rate > 5.0 °C/hr or spatial Z-score > 3.5σ vs MADIS median | Transducer resistance shift, loose terminal screw, or aspirated fan stoppage causing solar radiation trap | 4-wire bridge resistance measurement at junction box; certified reference thermometer comparison | IDW spatial median with 90% confidence band; raw stream preserved
Atmospheric Barometer Drop | Pressure (hPa) | Step > 6.0 hPa / 3hr or spatial residual > 5.0 hPa without rain/storm signals | Barometric port tubing dislocation, static head port blockage, or piezoresistive element rupture | Inspect port tube for condensation; zero-check with digital travelling reference barometer | Barometric height-adjusted regional consensus replacement
Calibration Drift Shift | Temperature / Pressure | Continuous progressive deviation > 0.5 unit/day accumulating persistent bias | Component aging, calibration degradation, or excitation voltage drift | Full multi-point laboratory calibration bath test | Time-decay weighted linear bias compensation flag
Frozen Sensor Flatline | All Parameters | Identical decimal reading across ≥ 5 consecutive observation cycles | A/D converter register lock, Modbus communication freeze, or mechanical float/vane stuck | Cycle logger DC power bus; verify sensor excitation current pulse | Impute from time-lagged autoregressive regional consensus
Humidity Boundary Lock | Relative Humidity (%) | Continuous reading stuck at 100.0% or 0.0% for ≥ 6 consecutive cycles | Water droplet accumulation on capacitive sensing element or ground short circuit | Inspect polymer sinter cap; clean with deionized water rinse and bake dry | Psychrometric reconstruction using dry-bulb temperature and regional dewpoint
Regional Weather Front | Temperature & Pressure | Multi-station simultaneous step change with neighbor agreement > 80% | Genuine meteorological boundary (cold front, sea breeze, thunderstorm gust front) | Verify radar reflectivity / satellite infrared cloud imagery | Preserve genuine atmospheric telemetry; suppress automated fault dispatch